## -summary-
File name as deployed	test.CR3
Program signature	52115
Deployment timestamp	2011-10-11 14:21:48 PST
Git repository tag	20111011_CFNT

Table name	Record interval	Notes
ts_data	100ms = 10hz
flux	30 min

StatCode	Description
Avg	arithmetic mean of samples over output period
Cov	covariance of samples over output period
Max	maximum value over output period
Med	median of samples over output period
Min	minimum value over output period
Smp	grab sample of instantaneous value
Smp+	grab sample of derived quantity
Std	standard deviation of samples over output period
Tot	total or cumulative sum of samples over output period
WV	Built-in wind-vector calculations

## ts_data
	Column label	Units	StatCode	Description	Notes
0	TIMESTAMP		--	end of period timestamp	as "YYYY-MM-DD hh:mm:ss.ss" (with quotes)
1	RECORD		--	sequential record number
2	Ux	m/s	Smp	horizontal wind component	(+) = into sonic array toward mounting bolt
3	Uy	m/s	Smp	horizontal wind component	cross wind speed, right-handed w.r.t. Ux/Uz
4	Uz	m/s	Smp	vertical wind speed	(+) = away from surface
5	Ts	C	Smp	sonic temperature	closely approximates virtual temperature
6	diag_sonic	arb	Smp	diagnostic code	bits 0-5 are warning flags; any value >0 has flags
7	co2_molar_density	umol/m^3	Smp	molar density of carbon dioxide
8	h2o_molar_density	mmol/m^3	Smp	molar density of water vapor
9	CO2	mg/m^3	Smp	mass conc. of carbon dioxide
10	H2O	g/m^3	Smp	mass conc. of water vapor
11	diag_irga	arb	Smp	diagnostic code	bits 0-20 are warning flags; any value >0 has flags
12	amb_tmpr	C	Smp	ambient temperature
13	amb_press	kPa	Smp	barometric (ambient) pressure
14	CO2_sig_strgth	arb	Smp	CO2 signal quality	nominal range 0-1, 1=good quality
15	H2O_sig_strgth	arb	Smp	H2O signal quality	nominal range 0-1, 1=good quality

## flux
	Column label	Units	StatCode	Description	Notes
0	TIMESTAMP		--	end of period timestamp	as "YYYY-MM-DD hh:mm:ss" (with quotes)
1	RECORD		--	sequential record number
2	l		Smp+	Obukhov length
3	Hs	W/m^2	Smp+	sensible heat flux	using sonic temp uncorrected for water vapor
4	tau	kg/(m s^2)	Smp+	momentum flux
5	u_star	m/s	Smp+	friction velocity
6	Ts_stdev	C	Smp+	std. dev. of sonic temperature
7	Ts_Ux_cov	C m/s	Smp+	covariance of sonic U,Ts
8	Ts_Uy_cov	C m/s	Smp+	covariance of sonic V,Ts
9	Ts_Uz_cov	C m/s	Smp+	covariance of sonic W,Ts
10	Ux_stdev	m/s	Smp+	std. dev. of sonic U component
11	Ux_Uy_cov	(m/s)^2	Smp+	covariance of sonic U,V
12	Ux_Uz_cov	(m/s)^2	Smp+	covariance of sonic U,W
13	Uy_stdev	m/s	Smp+	std. dev. of sonic V component
14	Uy_Uz_cov	(m/s)^2	Smp+	covariance of sonic V,W
15	Uz_stdev	m/s	Smp+	std. dev. of sonic W component	note no coordinate rotations are done
16	wnd_spd	m/s	WV	scalar mean horizontal wind speed
17	rslt_wnd_spd	m/s	WV	resultant mean horizontal wind speed
18	wnd_dir_sonic	degrees	WV	resultant mean WD relative to sonic
19	std_wnd_dir	degrees	WV	std. dev. of wind direction	via CampbellSci wind speed-weighted formula
20	wnd_dir_compass	degrees	WV	resultant mean wind direction	**see Note 1 below table
21	Ux_Avg	m/s	Avg	mean sonic axial component
22	Uy_Avg	m/s	Avg	mean sonic crosswind component
23	Uz_Avg	m/s	Avg	mean sonic vertical component
24	Ts_Avg	C	Avg	mean sonic temperature	closely approximates virtual temperature
25	sonic_azimuth	degrees	Smp	sample of sonic azimuth	user-defined; does not change
26	sonic_samples_Tot	samples	Tot	total recorded samples	should be 3000 for 5min, 18000 for 30min
27	no_sonic_head_Tot	samples	Tot	samples with no sonic signal
28	no_new_sonic_data_Tot	samples	Tot	samples with “NAN” sonic signal
29	amp_l_f_Tot	samples	Tot	low amplitude samples
30	amp_h_f_Tot	samples	Tot	high amplitude samples
31	sig_lck_f_Tot	samples	Tot	tracking errors
32	del_T_f_Tot	samples	Tot	speed of sound value warnings
33	aq_sig_f_Tot	samples	Tot	samples while re-acquiring signals
34	sonic_cal_err_f_Tot	samples	Tot	head calibration signature errors
35	Fc_wpl	mg/(m^2 s)	Smp+	CO2 flux	density (WPL) corrected
36	LE_wpl	W/m^2	Smp+	latent heat flux	density (WPL) corrected
37	Hc	W/m^2	Smp+	sensible heat flux	density (WPL) corrected
38	CO2_stdev	mg/m^3	Smp+	std. dev. of CO2 mass conc.
39	CO2_Ux_cov	mg/(m^2 s)	Smp+	covariance of sonic U & irga CO2
40	CO2_Uy_cov	mg/(m^2 s)	Smp+	covariance of sonic V & irga CO2
41	CO2_Uz_cov	mg/(m^2 s)	Smp+	covariance of sonic W & irga CO2
42	H2O_stdev	g/m^3	Smp+	std. dev. of H2O molar conc.
43	H2O_Ux_cov	g/(m^2 s)	Smp+	covariance of sonic U & irga H2O
44	H2O_Uy_cov	g/(m^2 s)	Smp+	covariance of sonic V & irga H2O
45	H2O_Uz_cov	g/(m^2 s)	Smp+	covariance of sonic W & irga H2O
46	Tc_stdev	C	Smp+	std. dev. of 'corrected' sonic temp. Tc
47	Tc_Ux_cov	C m/s	Smp+	covariance of sonic U,Tc
48	Tc_Uy_cov	C m/s	Smp+	covariance of sonic V,Tc
49	Tc_Uz_cov	C m/s	Smp+	covariance of sonic W,Tc
50	CO2_mean	mg/m^3	Smp+	mean CO2 mass concentration
51	H2O_mean	g/m^3	Smp+	mean H2O mass conc.
52	amb_press_mean	kPa	Smp+	mean barometric pressure
53	Tc_mean	C	Smp+	mean 'corrected' sonic temp.	**see Note 2 below table
54	rho_a_mean	kg/m^3	Smp+	mean ambient (moist) air density
55	Fc_irga	mg/(m^2 s)	Smp+	CO2 flux, uncorrected
56	LE_irga	W/m^2	Smp+	latent heat flux, uncorrected
57	CO2_wpl_LE	mg/(m^2 s)	Smp+	CO2 flux correction: WPL latent heat term
58	CO2_wpl_H	mg/(m^2 s)	Smp+	CO2 flux correction: WPL sensible heat term
59	H2O_wpl_LE	W/m^2	Smp+	LE flux correction: WPL latent heat term
60	H2O_wpl_H	W/m^2	Smp+	LE flux correction: WPL sensible heat term
61	irga_samples_Tot	samples	Tot	total recorded samples
62	no_irga_head_Tot	samples	Tot	samples with no irga signal
63	no_new_irga_data_Tot	samples	Tot	samples with “NAN” irga signal
64	irga_bad_data_f_Tot	samples	Tot	samples with suspect data	true for any diagnostic flag
65	gen_sys_fault_f_Tot	samples	Tot	general system faults
66	sys_startup_f_Tot	samples	Tot	samples while system is starting up
67	motor_spd_f_Tot	samples	Tot	samples while motor speed is out of bounds
68	tec_tmpr_f_Tot	samples	Tot	samples while detector temp is out of bounds
69	src_pwr_f_Tot	samples	Tot	samples while lamp power is out of bounds
70	src_tmpr_f_Tot	samples	Tot	samples while lamp temp is out of bounds
71	src_curr_f_Tot	samples	Tot	samples while lamp current is out of bounds
72	irga_off_f_Tot	samples	Tot	samples with gas head unpowered
73	irga_sync_f_Tot	samples	Tot	samples with channel sync error
74	amb_tmpr_f_Tot	samples	Tot	samples while ambient temp is out of bounds
75	amb_press_f_Tot	samples	Tot	samples while ambient pressure is out of bounds
76	CO2_I_f_Tot	samples	Tot	samples with CO2 intensity out of bounds
77	CO2_Io_f_Tot	samples	Tot	samples with CO2 reference out of bounds
78	H2O_I_f_Tot	samples	Tot	samples with H2O intensity out of bounds
79	H2O_Io_f_Tot	samples	Tot	samples with H2O reference out of bounds
80	CO2_Io_var_f_Tot	samples	Tot	samples with CO2 reference noise out of bounds
81	H2O_Io_var_f_Tot	samples	Tot	samples with H2O reference noise out of bounds
82	CO2_sig_strgth_f_Tot	samples	Tot	samples with CO2 reference ratio out of bounds
83	H2O_sig_strgth_f_Tot	samples	Tot	samples with H2O reference ratio out of bounds
84	CO2_sig_strgth_mean	arb	Avg	mean CO2 signal strength
85	H2O_sig_strgth_mean	arb	Avg	mean H2O signal strength
86	T_hmp_mean	C	Smp+	mean ambient temperature
87	e_hmp_mean	kPa	Smp+	mean derived H2O vapor pressure
88	e_sat_hmp_mean	kPa	Smp+	mean derived H2O saturation VP
89	H2O_hmp_mean	g/m^3	Smp+	mean derived H2O mass concentration
90	RH_hmp_mean	%	Smp+	mean relative humidity	derived from actual/sat. vapor pressure means
91	rho_a_mean_hmp	kg/m^3	Smp+	mean derived density of ambient (moist) air
92	Rn_Avg	W/m^2	Avg	mean net radiation	corrected for horizontal wind effects
93	Rn_meas_Avg	W/m^2	Avg	mean net radiation
94	par_totflx_Tot	mmol/m^2	Tot	total PAR flux density	cumulative per output interval
95	par_flxdens_Avg	umol/s/m^2	Avg	average PAR flux
96	034b_ws	m/s	WV	mean horizontal wind speed	comparable to wnd_spd from sonic
97	034b_wd	degrees	WV	unit vector mean wind direction	not comparable to sonic
98	034b_stdwd	degrees	WV	std. dev. of wind direction	Yamartino algorithm; not sonic-comparable
99	Rain_mm_Tot	mm	Tot	total precipitation	valid for rain, not snow/etc.
100	latitude_a	degrees	Smp	latitude degrees	(+ = North; - = South)
101	latitude_b	minutes	Smp	latitude decimal minutes	must combine with latitude degrees
102	longitude_a	degrees	Smp	longitude degrees	(+ = East; - = West)
103	longitude_b	minutes	Smp	longitude decimal minutes	must combine with longitude degrees
104	speed	m/s	Smp	ground speed
105	course	degrees	Smp	ground course bearing
106	magnetic_variation	unitless	Smp	magnetic variation	in degrees, easterly=(+); MagN+MagVar = TrueN
107	fix_quality	unitless	Smp	fix quality	0=invalid, 1=GPS, 2=diff. GPS, 6=estimated
108	nmbr_satellites	unitless	Smp	satellites used for fix
109	altitude	m	Smp	antenna altitude (geoidal height)	Earth datum: WGS 84
110	pps	ms	Smp	elapsed ms since last PPS signal from GPS
111	dt_since_gprmc	s	Smp	time since last GPRMC string	typically <1sec
112	gps_ready	unitless	Smp	GPS readiness	values less than 10 indicate poor GPS signal
113	max_clock_change	ms	Smp	maximum value clock was changed	values above 10 could indicate issues with the clock
114	nmbr_clock_change	samples	Smp	number of times clock was changed	cumulative value since program start
115	panel_tmpr_Avg	C	Avg	mean data logger panel temperature
116	batt_volt_Avg	V	Avg	mean data logger power source voltage
117	slowsequence_Tot	samples	Tot	total number of slow (1hz) scans performed

	**Note 1: not for straight-line Gaussian dispersion, OK for transport direction in variable-trajectory models (CRBasic Editor Help)
	**Note 2: ambient temp back-calculated using H2O conc. and virtual temp (which is approximated as sonic temp)